EPICARE BACKEND

AI-Driven Seizure Monitoring Integration Report

Prepared for: Client technical handoff

Scope: Backend API, AI decision pipeline, IoT Bridge, and Proteus/Arduino integration

Delivery date: 23 June 2026

Status: Implementation complete; .NET 8 compilation verification pending on an SDK-enabled machine

# Executive summary

The monitoring flow has been revised so the AI model and raw sensor signals are the default source of truth. Proteus no longer needs to send a preselected Normal, Warning, or Seizure answer during normal operation. A separate demo override remains available for deterministic presentations.

# Updated architecture

| Step | Component | Responsibility |
| --- | --- | --- |
| 1 | Proteus / Arduino | Sends raw EEG, ECG, EMG, and ACC readings. |
| 2 | IoT Bridge | Forwards readings to the backend and returns device commands when write-back is enabled. |
| 3 | Backend API | Builds the model input window and calls the AI prediction endpoint. |
| 4 | Decision Service | Combines AI risk with calibrated movement detection. |
| 5 | Arduino outputs | Receives N, P, or S to drive the LED and motor behavior. |

# Decision rules

| Input condition | Result | Device command |
| --- | --- | --- |
| AI result is Normal and below the configured risk threshold | Normal | N |
| AI returns final_prediction=1, Warning/Preictal, or exceeds the risk threshold | High risk | P |
| AI detects high risk and ACC movement exceeds the calibrated threshold | Seizure detected | S |
| AI returns an explicit Seizure or Ictal label | Seizure detected | S |
| DemoMode=true and Proteus supplies P or S | Forced demo state | P or S |

# Safety and integrity changes

Simulator isolation. Embedded Proteus state values are ignored when DemoMode is disabled.

Fallback isolation. AI HTTP failures now return a neutral unavailable result and do not copy the simulator answer into the AI result.

Movement validation. The presence of an ACC array alone no longer triggers a seizure. Horizontal movement and vertical deviation from the configured gravity baseline are evaluated against a threshold.

Raw data output. The Arduino sketch sends ACC data in all states and omits the internal state label by default, preventing label leakage.

# Configuration

| Setting | Default | Purpose |
| --- | --- | --- |
| Decision__DemoMode | false | Enables deterministic simulator overrides only when explicitly required. |
| Decision__AccGravityBaseline | 980 | Expected resting Z-axis value for the current Proteus sensor model. |
| Decision__AccVerticalScale | 50 | Normalizes vertical deviation in the movement score. |
| Decision__AccMovementThreshold | 2.5 | Minimum movement score required to escalate high risk to seizure. |
| AiModel__HighRiskThreshold | 0.7 | Probability threshold for a high-risk warning. |

# Delivered source code

| Area | Delivered content |
| --- | --- |
| Backend API | Complete EpiCare.Api source, configuration, deployment files, and solution metadata. |
| IoT Bridge | Serial-to-HTTP bridge with optional N/P/S command write-back. |
| Proteus / Arduino | Updated sketch with raw sensor output and optional demo-state transmission. |
| Documentation | Backend README and this technical integration report. |

# Validation status

The application configuration was parsed successfully as JSON, the patch passed git diff whitespace validation, and the Arduino sketch was checked to contain one active sensor-output function with demo-state transmission disabled by default. A full dotnet build could not be executed on the preparation machine because the .NET SDK was not installed in PATH. The client or deployment pipeline should run dotnet build EpiCare.Backend.sln using .NET 8 before release.

# Operational notes

ACC baseline and movement thresholds currently match the Proteus convention where resting Z is approximately 980. These values must be calibrated before connecting a physical accelerometer with different units. Monitoring history remains in memory and resets when the backend process restarts; persistent clinical history requires a database integration.

AI API endpoint update

The backend now targets the SeizeIT2 model API at https://seizeit2-api-production.up.railway.app with the prediction route POST /predict. Future model-server migrations should be handled by changing the environment variable AiModel__BaseUrl while keeping AiModel__PredictPath=/predict unless the new API contract changes.